DEVELOPMENT OF AN IN-VIVO ACOUSTIC DIAGNOSTIC FOR LITHOTRIPTER-INDUCED SHOCK-TISSUE INTERACTION (GR/N19243/01)

The following text is an abridged version of the final report to the sponsor (Engineering and Physical Sciences Research Council).

A fuller report will be made available from this web site when the appropriate permissions and authorisations have been obtained.

1.	Background

Extracorporeal shock wave lithotripsy (ESWL) is the preferred treatment for most kidney stones []. Several thousand shocks are focused onto the stone and, mainly through cavitation, break it down into fragments small enough to be passed naturally from the body (during urination), or dissolved with drugs []. Currently the shock wave focus is aligned with the stone at the start of the procedure using X-rays or ultrasound (although the images are not simple to interpret – Fig.1). Although the alignment might degrade during treatment as a result of patient or stone motion, subsequent checks of targeting accuracy are rarely possible because of X-ray exposure limitations. The first goal of this contract was to devise a passive instrument which could continually monitor this targeting. In addition, current practice exposes each stone to a fixed number of shocks (generated at a rate of 1-2 s-1), with no feedback available as to the degree of stone fragmentation. Hence the second goal of the passive device was to provide such feedback continuously. This would improve the procedure in a number of ways. If too few shocks are given to fragment the stone, the patient requires re-treatment, with commensurate costs in terms of finance, patient discomfort and waiting times, waiting lists, theatre time and resources (both facilities and staff). Currently >50% of patients require re-treatment. If the fixed number of shocks is greater than that required to fragment the stone sufficiently, this also wastes theatre and staff time. Furthermore, healthy tissue is unnecessarily exposed to the destructive shock waves, and the lithotripter shock wave source (which has a lifetime of a set number of shocks, and currently is replaced at a cost of thousands of pounds every few months) suffers unnecessary wear-and-tear [].

Fig. 1. X-ray image of a stone (and gas) before (left) and after (right) lithotripsy treatment.

Two schemes were devised to interpret the output of the device. First, an empirical scheme provided the surgeon with audio information from a loudspeaker in the theatre, so that he/she could ‘learn’ what sounds were characteristic of a successful treatment. Second, computation fluid dynamics (CFD) simulations were performed to improve understanding of how the signal detected by the device related to the conditions within the patient. The latter route would clearly have great potential for advancing other areas of science (and indeed has led to further grants in sonochemistry and erosion studies).

2. 	Key advances and supporting methodology

  - Key advances: Summary

In 1992-4, two of the investigators (Coleman and Leighton) correlated luminescent and acoustic emissions from the focus of a lithotripter to show that, under the influence of a lithotripter shock wave, a given bubble will emit two main shock waves, separated by about 200-300 (s [,]. Each shock wave signifies the end-point of a bubble collapse, the first being induced (at ~t=0 in Fig. 2) by the compressive shock wave from the lithotripter []. The bubble then rebounds, expands, and then collapses again, emitting a second shock wave (at ~t=220 (s in Fig. 2). When a multitude of bubbles follow this pattern, with a range of timings as result of their different locations and sizes, the characteristic ‘double-burst’ signal is detected (Fig. 3). One source of these shock waves is as follows. Cavitation bubbles close to the kidney stone involute on collapse, forming a jet. This jet may impact the stone directly or, if the bubble is further from the stone, will impact the far bubble wall, and a blast wave will be emitted. The device produced by this project operated by detecting the acoustic pressure signals generated by the cavitation. These differ depending on the accuracy of the targeting, and the degree of fragmentation of the stone.

Fig. 2. The top graph shows an idealised lithotripter pulse. On a longer common time axis the bubble radius (middle row) and the predicted pressure in the liquid at the bubble wall (bottom row) are shown, as calculated by the Gilmore-Akulichev model for a single bubble of initial radius 6 microns.

Fig. 3. A measured lithotripter pulse (upper plot) and a simultaneous output from a cavitation detector in vitro (lower plot). After the ‘first burst' of cavitation emissions (which can be compared in form to the later simulated ‘first burst’ in Fig. 12), the bubbles rebound. Then, after a ~300 (s prolonged expansion phase, collapse to generate the second burst.

The goal of the project was to build a passive device which could continuously monitor these signals and interpret them in terms of the integrity of the ESWL targeting and the degree of stone fragmentation. The advances through the stages of the development of the device are as follows. For conciseness, the in vitro (§2.1(a)), CFD (§2.1(b)) and clinical (§2.1(c)) components will be described separately, although in fact the interaction between the two was a major feature of the project.

2.1(a) In vitro experimental work

- A survey of the relevant in vitro and in vivo data was conducted and, by comparison with the Gilmore-Akulichev model for a spherical bubble, interpreted to estimate the likely initial radii of bubbles subjected to lithotripter pulses []. This key stage was vital, since the CFD work (see below) required knowledge of that initial bubble size. This provided the insight that the final stages of bubble oscillation were sufficiently low amplitude to allow their frequency to be related to bubble size (which would not have been appropriate during the earlier high-amplitude stages of oscillation, prior to the majority of dissipative losses). Use of the Gilmore-Akulichev spherical bubble model was also more relevant at this stage than the jetting stages (see §2.2(b), below). Since pulses occur at intervals of several hundred milliseconds, the bubble size at the end of a preceding lithotripter exposure would provide the appropriate initial bubble size for most exposures, rather than the naturally-occurring in vivo ambient bubble population (which is unknown, and would only be relevant at the outset of treatment).

- A previously non-operational lithotripter was commissioned for use in laboratory in vitro tests.

- A series of staged hydrophone measurements in vitro, with and without a stone present, allowed characterisation (in terms of frequency, amplitude and timing) of the basic characteristics of the emissions upon which the sensor prototypes would be based [].

- The results of a(iii) allowed further in vitro tests to compare hydrophone measurements with the signals from a prototype device built by the National Physical Laboratory to measure cavitation in industrial situations [8] (Fig. 4a). This collaboration benefited both organisations (the NPL system was a useful one to study, but would have been inappropriate for in vivo use).

- The results of a(iii) and a(iv) allowed the collaborative design of a prototype device with its manufacturers, Precision Acoustics Ltd. (PAL). Subsequent in vitro and in vivo testing of this, and a second refined prototype, allowed development of the third, and final, prototype (Fig. 4) [].

| Fig. 4. The NPL sensor (left) and the refined Precision Acoustics prototype (right). |  |  |
| --- | --- | --- |

2.1(b) Simulation and modelling

- A two-dimensional Free-Lagrange code for bubble collapse was investigated [,] and then adapted for three-dimensional axi-symmetric cases [,].

- Near field pressures were predicted for lithotripter-like cavitation events without solids present.

- Near field pressures were predicted for lithotripter-like cavitation events with plane solid half-spaces present (the solids having real material parameters, rather than simply being represented by rigid boundaries) [].

- Near field pressures were predicted for lithotripter-like cavitation events with solids containing notches present [14].

- Related work [] predicted the stresses within an elastic-plastic solid half-space close to a cavitation bubble. In addition to mapping these (and their evolution in time), an index indicating approach to material failure and its evolution in the solid was plotted [14].

- An initial estimation of the associated far-field pressures was made using the Helmholtz-Kirchoff method. Although this assumes linear propagation of the emissions, and hence at this stage contained engineering approximations whose effects had not been quantified, nevertheless the agreement with clinical measurements was extremely good [20].

- The far-field calculations were extended to cover the acoustic emissions from populations of thousands of bubbles, distributed throughout the cigar-shaped focus of the lithotripter field [14].

- The initial approximation of b(vi) was improved by using the Ffowcs-Williams-Hawkings method [14].

- Recognising that the procedure of b(viii) cannot describe any nonlinearities in the propagation, the size of  these were estimated using nonlinear propagation theory to examine propagation from the CFD near-field domain to the measurement point of the clinical sensor [14].

- A nonlinear propagation theory was developed for transient acoustic waves in bubbly water such as might be found near the lithotripter focus [,,].

- The work has been extended from single bubbles to predict at far field the emission from thousands of bubbles, distributed throughout the cigar-shaped focus of the lithotripter. This has produced new insights, particularly into how the duration of the bursts (d1 and d2 in Fig. 6) relate to the position of the sensor on the torso and the activity within the body.

2.1(c) Clinical work

- In the empirical tests (whereby a surgeon is supplied with an audio-frequency representation of the device output and becomes accustomed to which signals typify good and bad outcomes of ESWL) the three clinicians scored 100% in correctly identifying the features which would pertain to a clinical evaluation of the success of the treatment.

- The in vitro tests of a(iii)-a(v), particularly with the final prototype, enabled a range of characteristics to be identified for testing in the clinical environment, to determine the extent to which they indicate targeting and stone fragmentation [9,]. These were compared with the predictions of CFD and the level of agreement was very encouraging (compare ‘first burst; in Fig. 3 with Fig. 12; and see §2.2(c) for comparison of Figs. 8, 11 & 14) []. This enabled objective criteria to be drawn up, upon which the interpretive protocols on the final device are based.

|  | Fig. 5. Experimental set-up for in vitro tests involving ‘ping-pong’ ball target holder (see §2.2(a)). |
| --- | --- |

References

# Endnotes

Leighton T G.  From sea to surgeries, from babbling brooks to baby scans: bubble acoustics at ISVR. Proceedings of the Institute of Acoustics, 26, Part 1, 357-381, 2004, 357-381 (Tyndall Medal paper)

Leighton TG. The acoustics of gas bubbles in liquids. Anglo-French Physical Acoustics Group Meeting, (Institute of Physics), 2001, p. 7 (Keynote paper for Inaugural Meeting of this Group)

Leighton T G, Coleman AJ, Fedele F & White P R, A passive acoustic system for evaluating the in vivo performance of extracorporeal shock wave lithotripsy. UK Patent No. 0319863.7

Coleman A J, Choi M J, Saunders J E & Leighton T G, Acoustic emission and sonoluminescence due to cavitation at the beam focus of an electrohydraulic shock wave lithotripter, Ultrasound in Medicine and Biology, 18, 1992, 267-281

Coleman A J, Whitlock M, Leighton T G & Saunders J E, The spatial distribution of cavitation induced  acoustic emission, sonoluminescence and cell lysis in the field of a shock wave lithotripter, Phys. Med. Biol., 38, 1993, 1545-60

Leighton TG. Characterization of measures of reference acoustic cavitation (COMORAC). J. Acoust. Soc. Am., 108(5), Part 2, 2000, 2516. (Invited paper)

Cunningham K B, Coleman A, Leighton T G & White P R. Characterising in vivo acoustic cavitation during lithotripsy with time-frequency methods. Acoustics Bulletin 26(5), 2001, 10-16

Fedele F, Coleman AJ & Leighton TG. Use of a cylindrical PVdF hydrophone in a study of cavitation adjacent to stone phantoms during extracorporeal shockwave lithotripsy. Proc of the 9th Annual National Conference of the Institute of Physics and Engineering in Medicine,  (IPEM,  Bath) p. 66, 2003

Fedele F, Coleman AJ, Leighton TG, White PR & Hurrell AM. A new diagnostic sensor for Extra-corporeal Shock-Wave Lithotripsy. Acoustics Bulletin,  29, 2004, 34-39

Ball GJ, Howell BP, Leighton TG & Schofield MJ, Shock-induced collapse of cylindrical air cavity in water: Free-Lagrange simulation. Shock Waves, 10, 2000, 265-76

Ball GJ, Howell BP, Leighton TG & Schofield MJ, Shock-induced collapse of a cylindrical air cavity in water: a Free-Lagrange simulation. Proc. 22nd Int. Symposium on Shock Waves, London, July 1999, Paper 0060, published April 2000, 1363-1368

Jamaluddin AR, Ball GJ & Leighton TG. Free-Lagrange simulations of shock/bubble interaction in shock wave lithotripsy. The 24th International  Symposium on Shock Waves, Beijing, China, July 11 16, 2004 (in press)

Jamaluddin AR, Ball GJ & Leighton TG. Free-Lagrange simulations of shock/bubble interaction in shock wave lithotripsy. Proc. 2nd Int. Conf. on Computational Fluid Dynamics, ICCFD, Sydney, Australia, 15-19 July 2002, pp 541-546

Jamaluddin AR, Turangan CK, Ball GJ & Leighton TG. Free-Lagrange simulations of the jetting collapse of air bubbles with application to extracorporeal shock wave lithotripsy. Proceedings of the Royal Society A (in preparation)

Turangan CK & Ball GJ. A Free-Lagrange Simulation of Cavitation Bubble collase near a Rigid Boundary. The 23rd International Symposium on Shock Waves, Fort-Worth TX, USA, July 2001, pp793-799

Leighton TG, Meers SD & White PR. A nonlinear time-dependent inversion to obtain bubble populations from acoustic propagation characteristics. Proceedings of the Royal Society (in press; publ.Royal Society FirstCite Website May 2004)

Leighton TG & Dumbrell HA. New approaches to contrast agent modelling. Proceedings of the First Conference in Advanced Metrology for Ultrasound in Medicine (Journal of Physics Conference Series) (2004, accepted).

Leighton TG & Dumbrell HA. Modeling acoustic propagation through clouds of nonlinearly pulsating bubbles in vivo. Ultrasound Med.Biol. (2004, in preparation)

Fedele F, Coleman AJ, Leighton TG, White PR & Hurrell AM.  A new sensor for detecting & characterising acoustic cavitation in vivo during ESWL. Proceedings of the Institute of Acoustics, 26, Part 1, 2004, 422-432

Leighton TG ‘From babbling brooks to baby scans, from seas to surgeries: The pressure fields produced by non-interacting spherical bubbles at low & medium amplitudes of pulsation’, International Journal of Modern Physics B (in preparation) 2004